|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **SOLUTION LOT#:** |  | **SOLUTION EXP.** (fr. Bottle)**:** |  | **SOLUTION EXP.** (open + 90 days)**:**  |  |
| **HLD CHANGE DATE\*:** |  | **HLD EXPIRY** (dispense + 21 days)**:** |  | **STAFF INITIALS** (dispense + test)**:** |  |

|  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **DATE + TIME**(yyyy-mm-dd: hh:mm) | **PATIENT or MRN#:** | **PROBE#** | **TECH** | **PROBE CONDITION GOOD?YES: (✓) / NO: specify** | **TEST STRIP (test EACH use) (✓) :** | **SOLUTION TEMP. (≥ 20˚C):** | **SOAK/CONTACT TIME** (= 5 mins): | **DISINFECTED BY (initials):** |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
| **DATE + TIME**(yyyy-mm-dd: hh:mm) | **PATIENT or MRN#:** | **PROBE#** | **TECH** | **PROBE CONDITION GOOD?YES: (✓) / NO: specify** | **TEST STRIP (test EACH use) (✓) :** | **SOLUTION TEMP. (≥ 20˚C):** | **SOAK/CONTACT TIME** (= 5 mins): | **DISINFECTED BY (initials):** |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |

* *\* Please start new log sheet for solution change (as it is necessary to record new dispense dates).*
* *After opening, the solution may be stored for use for up to 90 days (provided the 90 days do not extend past the expiration date on the container). When poured into a secondary container (i.e., basin), Revital-Ox RESERT HLD solution can be reused for up to 21 days or until its concentration falls below the minimum recommended concentration of 1.5% hydrogen peroxide. The MRC must be monitored before each use using the Revital-Ox™ Resert® R60 Solution Test Strip.*

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **TEST STRIP NAME** | **LOT #** | **VALIDATE w FRESH SOLUTION (✓)** | **STRIP EXPIRY DATE (fr. Bottle)** | **TEST STRIP EXP. (open + 180 days)** | **INITIALS** |
| Revital-Ox RESERT R60 Solution Test Strip |  |  |  |
|  |  |  |  |  |  |
|  |  |  |  |  |  |

*The shelf life after the first opening of the test strip bottle is 180 days, or the labeled expiration date, whichever is sooner. Do not use this product beyond the expiration date.*

**ALWAYS USE HLD SOLUTION ABOVE 20˚C. TEST STRIP + DISINFECTION MAY NOT BE EFFECTIVE IF USED WITH HLD SOLUTION BELOW 20˚C.**